## Own Method
ID_Institute | Name_Institute | Instrument manufacturer | Model type and number | Analytical column | column dimensions | solvent used for reconstitution | injection volume | column temperature | composition mobile phase | mobile phase gradient programme | flow rate | scan range | resolving power with referenced m/z | accuracy MS | accuracy MS/MS | width of the ion | Ionization | methods of mass fragmentation | Collision energy | software target | procedure confirmation targets | software/databases used for suspect screening | Procedure for confirmation of suspects and unknowns | rt prediction method | fragmentation prediction method | Unit
1 | Aarhus University | Thermo Fisher Scientific | Q Exactive HF | PepMap RSLC, C18 | 750; 0.75; 2 | MeOH/Water 4.8/95.2 v/v | 5 | 50 | A=2% acetontrile+0.1% formic acid+water; B=2% water+0.1% formic acid+acetonitrile | 10% B(0), 10%-95% B over 70 min; 95%B (10 min), 95% - 10%B over 1.5 min; 10%B (18.5) - total 100 min | 0.0002 | 90-900 | 60000 @m/z 200; 240000 @m/z 200 (additional method) |  |  |  | ESI (P) | HCD | 35 | Compound Discoverer | MS, MS/MS, Library comparison (mzCloud) | Compound Discoverer | MS, MS/MS, Library comparison (mzCloud, In house, HighRes Lib., ChemSpider with in silico comparison) | none applied | mzLogic, FISH | na
2 | Athens University | Bruker | UHPLC-QTOF-MS (Maxis Impact) | Acclaim RSLC C18 column + guard column, ACQUITY UPLC BEH C18 1.7 μm, VanGuard Pre-Column | 100; 2.1; 2.2 | Methanol/Water 50/50 | 5 | 30 | A=H2O:MeOH (90:10) + 5 mM ammonium formate + 0.01% formic acid; B=MeOH + 5 mM ammonium formate + 0.01% formic acid | 1% B at 0-1 min; 39% B at 3 min; 99.9% B at 14 min and held until 16 min, 1% B at 16.1-20 min | 200 µL/min at 0-3 min, 400 µL/min at 14 min, 480 µL/min at 16-19 min, 200 µL/min at 19.1-20 min | 50-1000 | 24654 @ m/z 226.9515; 16948 @ m/z 430.9137; 21562 @ m/z 702.8636 | 3 mDa | 5 mDa |  | ESI (P) | CID | 4 eV/ 25 eV | Target Screener R package | MS, RT, MS/MS | AutoNon-Target R package, Norman SusDat | MS, RT, MS/MS, prediction of retention time, Exp. RTI match, Molecular Formula evaluation | QSRR (DOI: 10.1016/j.jhazmat.2018.09.047) | MetFrag, FragPred, and CMF-ID | ug/L
3 | BRGM | Waters | Acquity UPLC I Class - Xevo G2S Q-TOF | Waters BEH C18 | 150; 2.1; 1.8 |  | 2 (2 samples) or 5 (3 samples) | 35 | A=0.05 % formic acid in water; B = 0.05 % formic acid in acetonitrile | 2 % B (0.5); 99 % B (18); 99 % B (22); 2 % B (25) | 0.45 | 50-1200 | 35,000 @m/z 550 | 0.4 ppm | 0.4 ppm |  | ESI (P) | CID | Ramp 6 to 45eV | UNIFI | MS, RT, MS/MS | Homemade database | MS, RT, MS/MS |  | UNIFI
4 | CNR-IRSA Italy | Sciex | 5600+ TripleTOF | Zorbax Eclipse Plus C18 + Waters BEH Shield C18 | 5; 2.1; 1.8 + 150; 2.1; 1.7 | Methanol/Water 10/90 | 5 | 40 | A=water+0.1% formic acid; B=acetonitrile+0.1% formic acid | 2% B at 0 min; 2% B at 0.5 min; 20% B at 1.5 min; 100% B at 16 min; 100% B at 24.5 min; 2% B at 24.7; 2% B at 35 min | 0.2 | 60-900 | 25,000 @m/z 200; ≥30,000 @m/z 900 | 3ppm | 3ppm |  | ESI (P) | CID | 35 eV (CES 15 eV) | SCIEX OS, PeakView | MS, RT, MS/MS | Digital Sample Freezing Platform, SCIEX OS, PeakView | MS, MS/MS, prediction of retention time | Digital Sample Freezing Platform |  | -
5 | CNR-IRSA Italy-Brugherio
6 | EAWAG | Thermo Fisher Scientific | Orbitrap Lumos Fusion | Waters Atlantis C18 | 100; 3.0; 3.0 | injected as is | 5 | 30 | A=water+0.1% formic acid; B=methanol+0.1% formic acid | 0%B(for 1.5 min);0%B- 95%B over 17min;95%B for 10min; 95%B-0%B over 0.5min; 0%B for 4min | 0.3 | 100-900 | MS1: 240K @ m/z 200; MS2: 15K @m/z 200 | < 5 ppm | < 5 ppm |  | ESI (P) | HCD | 15, 50, 100 |  |  | CompoundDiscoverer 3.1 | MS, RT, MS/MS libraries, in silico fragmention MS2, prediction of retention time | http://rti.chem.uoa.gr/ | MassFrontier
7 | EHU | Thermo-Fisher Scientific, CA, USA | Thermo Scientific Dionex UltiMate 3000 UHPLC coupled to a Thermo Scientific Q Exactive Focus q-Orbitrap mass spectrometer. | ACE UltraCore 2.5 SuperPhenylHexyl | 100; 2.1; 2.5 | Methanol | 5 | 35 | A=water+0.1% formic acid; B=acetonitrile+0.1% formic acid | 15 %B(0);15 %B-30% over 4min;30 %B-50% over 4min;50%B(12);50 %B-90% over 10min;90%B(5);90%B-15% over 3min;15%(5) | 0.3 | 70-1050 | 70,000 @m/z 200 (MS1); 17.500 @m/z 200 (MS2) |  |  | 3 m/z | ESI (P) | HCD | 30 eV | Trace Finder |  | Compound Discoverer v2.1 / Stoffident, MZCloud | MS, RT, MS/MS, prediction of retention time |  | MetFrag
8 | ICRA | Thermo Fisher Scientific | Orbitrap Velos | ZORBAX Eclipse XDB-C18 | 150; 4.6; 5.0 |  | 10 | 30 | A: Water 10 mM ammonium formate pH3; B=Acetonitril | 95/5 at 0 min; 95/5 at 1 min; 5/95 at 10 min; 0/100 at 13 min; 0/100 at 15 min; 95/5 at 16 min; 95/5 at 17 min | 0.5 | 60-900 | 60,000 and 30,000 @m/z 400; | 5 ppm | 5 ppm |  | ESI (P) | CID | 35 eV | Xcalibur | MS, RT, MS/MS | DSFP, MassBank | MS, RT, MS/MS, prediction of retention time | no | MetFrag
9 | INERIS | Agilent | QTOF 6550 | Zorbax SB AQ | 150; 2.1; 1.8 | as provided | 5 | 40 | A=water+ 1 MM ammonium acetate; B=methanol | T0min : 100/0, T2min : 100/0, T14min : 2/98, T20min : 2/98, T21min : 100/0, T24min100/0 | 0.4 | 50-1000 | 10,000 @m/z 50; 30,000 @m/z 1000 | 5 ppm | 10 ppm |  | ESI (P) | CID | 20 and 40 eV |  |  | Masshunter | MS, RT, MS/MS
10 | Irstea | Waters | UPLC H-Class / Xevo-G2S-QTOF | Waters HSS T3 | 100; 2.1; 1.8 | - | 5 | 30 | A=water + 0,1% formic acid ; B=acetonitrile + 0,1% formic acid | 2%B(0);2%B(2);2%-99%B over 13min; 99%B(17); 99%-2%B over 1min; 2%B(20) | 0.5 | 50-950 | 24,000 @m/z 556,2771 | 0.3 ppm |  |  | ESI (P) | CID | 10-45 eV | Mass Lynx |  | Unifi | MS, RT, MS/MS, prediction of retention time | lab made correlation
11 | ISA | Bruker Daltonics | Maxis Plus | Acclaim RSLC 120 C18 | 100; 2.1; 2.2 |  | 5 | 30 | A=water/MeOH 90/10+5mM ammonium formate +0,01% Formic Acid ; B=MeOH+5mM ammonium formate+0,01% Formic Acid | 1%B (0,1'); 1% B-40%over 3'; 40%B-99% over 13'; 99%B (2) | 0.4 | 60-900 | 25,000@m/z430; | 5 ppm | 10ppm |  | ESI (P) | CID | 40eV | TASQ |  | TASQ with PesticideScreener and ToxScreener | MS, RT, MS/MS, isotope pattern |  | Compound Crawler, Met-Frag
12 | LEESU | Waters | Vion IMS Q-TOF | Acquity UPLC BEH C18 | 100; 2.1; 1.7 | methanol/water 20/80 | 5 | 40 | A=water+0.1% formic acid; B=Acetonitrile+0,1% formic acid | 2% B (0); 2%B (1); 2% B-98%B over24min; 98%B (26); 98%B-2% B over 1.5min; 2%B (34) | 0.45 | 60-900 | 32,000 @m/z 556 | 10 ppm |  |  | ESI (P) | CID | 10-45eV | UNIFI | MS, RT | UNIFI, in-house-library | MS, RT |  | In silico fragmentation with UNIFI (suspect only)
13 | Örebro University | Waters | G2-XS-qToF | Acquity UPLC BEH C18 | 100; 2.1; 1.7 |  | 5 | 50 | A=water+ 0.1 formic acid B=ACN + 0.1% formic acid | 5%B(0); 5%B-95%A over 0.5min; 100%B-0%A over 13min; 100%B-0%A over 14min; 5%B-95%A over 14.10; 5%B-95%A over 16min | 0.3 | 60-900 |  | 5 ppm | 5 ppm |  | ESI (P) |  | 6, 10-45 eV | UNIFI |  | UNIFI |  |  | MetFrag
14 | QAEHS | Sciex | X500R | Phenomenex Kinetex Biphenyl | 50; 2.1; 2.6 |  | 5 | 50 | A=water+0.1% formic acid; B=methanol+0.1% formic acid | 0%B for 0.3 min; 0-100% B over 18.7 min; 100 % B for 2.9 min; 100-0% B over 0.1 min; hold 0% B for another 3 min | 0.4 | 100-950 |  |  |  |  | ESI (P) | CID | 35 |  |  | SciexOS / NIST
15 | Rijkswaterstaat | Thermo Fisher Scientific | Orbitrap Qexactive HF | Waters Xbridge C18 | 150; 2.1; 3.5 | changed vial | 5 | 35 | A = 95% (4 mM NH4Formate in water+ 0.1% formic acid) /5% methanol B = 99.8% methanol / 0.2% (4 mM NH4Formate in water+ 0.1% formic acid) | 5%(0 min); 5%B(0.1 min); 5%B-16%B over 1 min, 16%B-98%B over 18 min;98%B(6 min); 98%B-5%B over 0.1 min; 5%B(35 min). | 0.3 | 60-900 | 60,000 @m/z 400; | 2 | 3 |  | ESI (P) | HCD | NCE35 |  |  | Compound Discoverer | MS, RT, MS/MS, prediction of retention time | log Kow correlation (Letzel) / Norman RTI model | Compound Discoverer | µg/l
16 | SLU | Waters | G2-S QTOF | Acquity HSS-T3 C18 | 100; 2.1; 1.8 |  | 10 | 40 | A=5mM ammonium formate (in water)+0.01% formic acid; B=acetonitrile+0.01% formic acid | 5%B for 0.5min(0-0.5); 5-95%B over 15.5min(0.5-16);95%-99%B in 0.1min(16-16.1);99% B for 2.9 min (16.1-19);99%-5%B in 0.1 min(19-19.1); 5%B for 1.9min (19.1-21) | 0.5 | 50-800 | 35,000 @m/z 300; 32,000 @m/z 550 | 2 mDa | 3mDa |  | ESI (P) | CID | 10-45 eV | UNIFI |  | UNIFI, MetFrag, MassBank |  |  | UNIFI, MetFrag
17 | University Jaume I | Waters | QTOF Xevo G2 | Cortecs C18 | 100; 2.1; 2.7 |  | 5 | 40 | A=water+0.01%HCOOH; Methanol+0.01%HCOOH | 10 %B-90% over 14 min(0-14); 90%B for 2 min (14-16); 90%B-10% over 0.1 min (16-16.1); 10%B for 1.9 min (16.1-18) | 0.3 | 50-1000 | 20,000 @m/z 556 | 5 ppm | 5 ppm |  | ESI (P) | CID | 15-40 eV | UNIFI | MS, RT, MS/MS | UNIFI
18 | University of South Bohemia | Thermo Fisher Scientific | Qexactive | Phenomenex Kinetex EVO C18 | 50; 2.1; 5 | as delivered | 5 | 23 | A=water+10mmol amonium acetate + 0.1% formic acid; B=methanol+10mmol amonium acetate +0.1% formic acid | 10 %B(0);10 %B-50% over 4min;50%B(10);50 %B-95% over 10min;95%B(10) | 0.3-0.4 | 100-1000 | 70000 FWHM @m/z 250; 40000 FWHM @m/z 800 | 2 ppm | 5 ppm |  | ESI (P) | HCD |  | CompoundDiscoverer | MS, MS/MS, izotope patterns | Compound Discoverer embeded + own library, myCloud, Chemspider with using MassBank | MS, MS/MS, prediction of MS/MS, RT if available |  | MassFrontier
19 | Vitens | Sciex | Triple TOF 6600+ | Waters Xselect HSS C18 | 150; 4.6; 3.5 | Acetonitril/Water 1/99 | 1000 | 35 | A=water+0.1% formic acid; B=Acetonitril | 1%B(0 min), 1%B (2 min), 99%B (33 min), 99%B (45 min), 1%B (47 min) | 1 | 77-1000 | 28,000 @m/z 130 43,000 @m/z 609 | 2 ppm | 5 ppm |  | ESI (P) |  | 35 eV | PeakView | MS, RT, MS/MS | Meta | MS, RT, MS/MS, prediction of retention time
20 | VITO | Thermo Fisher Scientific | Q Exactive | Acquity UPLC BEH C18 | 100; 2.1; 1.7 | direct injection | 5 | 50 | A=5mM NH4Ac in H2O; B= 5mM NH4Ac in MeOH | 5 %B(0 - 2 min); 5 %B - 90%B over 15min; 90%B (5 min); 5 %B (0.1 min); 5%B (4.9 min) | 0.35 | 60-900 | 70,000 @m/z 60; 70,000 @m/z 900 | 5 ppm | 5 ppm |  | ESI (P) | HCD | 30-40 eV | Trace Finder | MS, RT, MS/MS | Compound discoverer 3.1 | MS, RT, MS/MS, prediction of retention time, mzcloud |  | MetFrag | ug/L
21 | VU Amsterdam | Bruker daltonics | Compact QToF model (Bruker, Germany) | Lunar Omega | 100; 2.1; 3 | was already solved | 5 | 30 | A=Water; B=Methanol | 0% B(0); 0%B-80% over 17.78min;80% B for 1.78min;80%B-1% over 4.44min;1%B for 4.44min | 0.3 | 30-1300 | 25000 @m/z 1200 | 3 ppm | 5 ppm | 0.014 FWHM @ m/z 200 | ESI (P) | CID | 20eV + 40eV (stepping) | - | - | MetaboScape® 4.0 | MS, MS/MS

## Pre-defined method
ID_Institute | Name_Institute | Instrument manufacturer | Model type and number | Analytical column | column dimensions | Solvent used for sample reconstitution | injection volume | column temperature | composition mobile phase | mobile phase gradient programme | flow rate | scan range | resolving power with referenced m/z | accuracy MS | accuracy MS/MS | width of the ion | Ionization | methods of mass fragmentation | Collision energy | software target | procedure confirmation targets | software/databases used for suspect screening | Procedure for confirmation of suspects and unknowns | rt prediction method | fragmentation prediction method | Unit
1 | Aarhus University
2 | Athens University | Bruker | UHPLC-QTOF-MS (Maxis Impact) | ACQUITY UPLC BEH C18 (Waters) + VanGuard Pre-Column | 100; 2.1; 1.7 | AcN/Water 50/50 | 5 | 50 | A=water+5mM Ammonium Formate (adjusted to pH=3.0); B=ACN+0.1% formic acid | 13% B at 0-0.44 min; 50% B at 8.89 min; 95% B at 9.56 min and held until 10.89 min, 13% B at 11.11-13.33 min | 0.3 | 60-900 | 25584 @ m/z 226.9515; 24666 @ m/z 430.9138; 34047 @ m/z 702.8635 | 3 mDa | 5 mDa |  | ESI (P) | CID | 4 eV/ 35 eV | Target Screener R package | MS, RT, MS/MS | AutoNon-Target R package, Norman SusDat | MS, RT, MS/MS, prediction of retention time indices, Exp. RTI match, Molecular Formula evaluation | RTI (http://rti.chem.uoa.gr/) | MetFrag, FragPred, and CMF-ID | ug/L
3 | BRGM | Waters | Acquity UPLC I Class - Xevo G2S Q-TOF | Waters BEH C18 | 150; 2.1; 1.8 |  | 5 | 50 | A=5mM aqueous ammonium formate(pH3); B = 0.1 % formic acid in acetonitrile | 13 % B (0.5); 50 % B (10); 95 % B (10,75); 95 % B (12,25); 13% B (12,50); 13% B (15) | 0.4 | 60-900 | 35,000 @m/z 550 | 0.4 ppm | 0.4 ppm |  | ESI (P) | CID | Ramp 10 to 45eV | UNIFI | MS, RT, MS/MS | UNIFI, homemade database | MS, RT, MS/MS |  | UNIFI
4 | CNR-IRSA Italy | Sciex | 5600+ TripleTOF | Zorbax Eclipse Plus C18 + Waters BEH Shield C18 | 5; 2.1; 1.8 + 150; 2.1; 1.7 | Methanol/Water 10/90 | 5 | 40 | A=5mM aqueous ammonium formate, pH3 B=acetonitrile, 0.1% formic acid | 13% B at 0 min; 13% B at 0.5 min; 50% B at 10 min; 95% B at 10.75 min; 95% B at 12.25 min; 13% B at 12.50; 13% B at 15 min | 0.4 | 60-900 | 25,000 @m/z 200; ≥30,000 @m/z 900 | 3ppm | 3ppm |  | ESI (P) | CID | 35 eV (CES 15 eV) | SCIEX OS, PeakView | MS, RT, MS/MS | Digital Sample Freezing Platform, SCIEX OS, PeakView | MS, MS/MS, prediction of retention time | Digital Sample Freezing Platform |  | -
5 | CNR-IRSA Italy-Brugherio | Thermo Fisher Scientific | Q Exactive Focus | RESTEK Raptor C18 cod.9304262 | 15; 2.1; 1.80 | 1:1 with 5 mM aqueous ammonium formate, pH 3.0 | 10 | 30 | A=5 mM aqueous ammonium formate, pH 3.0; B=Acetonitrile+0.1% formic acid | 5 %B(0) for 0.2min; 5 %B-20% over 0.8min; 20%B-95 %B over 20min; 95%B for 5min; 5%B over 0.5 min; 5%B for 30 min; | 0.25 | 75-950 | 70,000 @m/z 500; Lock mass; | <4 ppm | < 4 ppm |  | ESI (P) | CID | 20-40 |  |  | CompoundDiscoverer 3.1 | MS isotope pattern, mzCloud library, ChemSpider, mass list search, in silico fragmention MS2, prediction of retention time (http://rti.chem.uoa.gr/) |  | CompoundDiscoverer 3.1
6 | EAWAG | Thermo Fisher Scientific | Orbitrap Lumos Fusion | Acquity UPLC BEH C18 | 150; 2.1; 1.7 | injected as is | 5 | 50 | A=water+5mM NH4FA,pH3; B=ACN+0.1% formic acid | 13%B for 0.5 min; 13%B-50%B over 9.5min; 50%B-95%B over 0.75min; 95%B for 1.5min; 95%B-13%B over 0.25min; 13%B for 2.5min | 0.4 | 100-900 | MS1: 60K @ m/z 200; MS2: 15K @m/z 200 | < 5 ppm | < 5 ppm |  | ESI (P) | HCD | 15, 50, 100 |  |  | CompoundDiscoverer 3.1 | MS, RT, MS/MS libraries, in silico fragmention MS2, prediction of retention time | http://rti.chem.uoa.gr/ | MassFrontier
7 | EHU | Thermo-Fisher Scientific, CA, USA | Thermo Scientific Dionex UltiMate 3000 UHPLC coupled to a Thermo Scientific Q Exactive Focus q-Orbitrap mass spectrometer. | Kinetex 1.7 um XB-C18 100A | 150; 2.1; 1.7 | Methanol | 5 | 50 | A=5 mM aqueous ammonium formate, adjusted to pH 3.0. with formic acid; B=acetonitrile+0.1% formic acid | 13 %B(0);13 %B(0.5);13 %B-50% over 9.5min;50 %B-90% over 4min;50%B(12);50 %B-95% over 0.75min;95%B(2);95%B-13% over 0.75min;13%(2.5) | 0.4 | 60-900 | 70,000 @m/z 200; 17.500 @m/z 200 |  |  | 3 m/z | ESI (P) | HCD | 10-30-70 eV | Trace Finder | MS, RT, MS/MS | Compound Discoverer / Stoffident, MZCloud | MS, RT, MS/MS, prediction of retention time |  | MetFrag
8 | ICRA | Thermo Fisher Scientific | Orbitrap Velos | Phenomenex Luna C18 | 100; 2.1; 1.6 |  | 10 | 40 | A=5mM ammonium formate (in water, pH=3); B=acetonitrile+0.1% formic acid | 87/13 at 0 min; 87/13 at 0.30; min50/50 at 5.90 min; 5/95 at 6.40 min; 5/95 at 7.30 min; 87/13 at 7.40 min; 87/13 at 8.90 min | 0.4 | 60-900 | 60,000 and 30,000 @m/z 400; | 5 ppm | 5 ppm |  | ESI (P) | CID | 35 eV | Xcalibur | MS, RT, MS/MS | DSFP, MassBank | MS, RT, MS/MS, prediction of retention time | no | MetFrag
9 | INERIS | Agilent | QTOF 6550 | Zorbax SB AQ | 150; 2.1; 1.8 |  | 5 | 50 | Mobile Phase A: 5 mM aqueous ammonium formate, adjusted to pH 3.0. Mobile Phase B: Acetonitrile containing 0.1% formic acid | as required | 0.4 | 60-900 | 10,000 @m/z 50; 30,000 @m/z 1000 | 5 ppm | 10 ppm |  | ESI (P) | CID | 20 and 40 eV |  |  | Masshunter | MS, RT, MS/MS
10 | Irstea | Waters | UPLC H-Class / Xevo-G2S-QTOF | Waters HSS T3 | 100; 2.1; 1.8 | - | 5 | 50 | A=5mMaqueous ammonium formate + 0,1% formic acid; B=acetonitrile + 0,1% formic acid | 13%B(0);13%B(0,1);13%-50%B over 1,9min; 50%-5%B over 5,2min; 5%B(8,2); 5%-87%B over 0,1min; 87%B(10) | 0.4 | 60-900 | 25,000 @m/z 556,2771 | 0.3 ppm |  |  | ESI (P) | CID | 10-45 eV | Mass Lynx |  | Unifi
11 | ISA | Bruker Daltonics | Maxis Plus | Agilent Eclipse XDB C18 | 150; 2,1; 1,8 |  | 5 | 50 | A= 5 mM aqueous ammonium formate, adjusted to pH 3.0.; B=Acetonitrile+ 0.1% formic acid | 13%B(0);13%B-50% over 10min;50 %B-95% over 0,75min; 95%(B) 1,5min; 95%-13%(B) over 0,5min; 13%B(2,5) | 0.4 | 60-900 | 25,000@m/z430; | 5ppm | 10ppm |  | ESI (P) | CID | 40 eV | TASQ |  | TASQ with PesticideScreener and ToxScreener | MS, RT, MS/MS, isotope pattern |  | Compound Crawler, Met-Frag
12 | LEESU | Waters | Vion IMS Q-TOF | Acquity UPLC HSS C18 | 150; 2.1; 1.8 |  | 5 | 50 | A=water+ammonium formate 5mM pH 3,1; B=Acetonitrile+0,1% formic acid | 13% B (0); 13%B (0,5); 13% B-50%B over 9,5min; 50%B-95%B over 0,75min; 95%B (12,25); 95%B - 13%B over 0,25min; 13%B (15) | 0.4 | 60-900 | 32,000 @m/z 556 | 10 ppm |  |  | ESI (P) | CID | 10-45eV | UNIFI | MS, RT | UNIFI, in-house-library | MS, RT |  | In silico fragmentation with UNIFI (suspect only)
13 | Örebro University | Waters | G2-XS-qToF | Acquity UPLC BEH C18 | 100; 2.1; 1.7 |  | 10; 5 (blank) | 30 | A=water+5mM ammonium formate B=ACN + 0.1% formic acid | 13%B(0); 13%B-87%A over 0.33min; 50%B-50%A over 2.32min; 95%B-5%A over 7.17min; 95%B-5%A over 8.17: 13%B-87%A over 8.33min; 87B-13%A over 10min | 0.4 | 60-900 |  | 5 ppm | 5 ppm |  | ESI (P) |  | 6, 10-45 eV | UNIFI |  | UNIFI |  |  | MetFrag
14 | QAEHS | Sciex | X500R | Phenomenex Luna C18 | 150; 2.0; 3 |  | 5 | 40 | as requested | as requested | 0.24 | as requested |  |  |  |  | ESI (P) | CID | 35 |  |  | SciexOS / NIST
15 | Rijkswaterstaat | Thermo Fisher Scientific | Orbitrap Qexactive HF | Waters Xbridge C18 | 150; 2.1; 3.5 |  | 5 | 50 | A=5 mM NH4Formate buffer in LC water to pH3 with formic acid; B = 100% Acetonitril+0.1% Formic acid | Compulsory Grad NT2019 after Method/column translation: 13%B(0 min);13%B(0.97 min);13%B-50%B over 18.47 min; 50%B-95%B over 1.46 min; 95%B (7.1 min); 95%B-13%B over 0.5 min; 13%B (6.5 min) | 0.21 | 60-900 | 60,000 @m/z 400; | 2 | 3 |  | ESI (P) | HCD | NCE35 |  |  | Compound Discoverer | MS, RT, MS/MS, prediction of retention time | log Kow correlation (Letzel) / Norman RTI model | Compound Discoverer | µg/l
16 | SLU | Waters | G2-S QTOF | Acquity HSS-T3 C18 | 100; 2.1; 1.8 |  | 5 | 40 | A=5mM ammonium formate (in water, pH=3); B=acetonitrile+0.1% formic acid | 13%B for 0.3min(0-0.3); 13-50%B over 6.4min(0.3-6.7);50%-95%B in 0.5min(6.7-7.2);95% B for 1min (7.2-8.2);95%-13%B in 0.1 min(8.2-8.3); 13%B for 1.7min (8.3-10) | 0.4 | 60-900 | 35,000 @m/z 300; 32,000 @m/z 550 | 2 mDa | 3mDa |  | ESI (P) | CID | 10-45 eV | UNIFI |  | UNIFI, MetFrag, MassBank |  |  | UNIFI, MetFrag
17 | University Jaume I | Waters | QTOF Xevo G2 | Cortecs C18 | 150; 2.1; 2.7 |  | 5 | 40 | A=water+5mM NH4COOH (at pH3); Acetonitril+0.1%HCOOH | 13 %B for 0.67 min(0-0.67);13 %B-50% over 12.66 min (0.67-13.33);50%B-95% over 1 min(13.33-14.33);95%B for 2 min (14.33-16.33); 95%B-13% over 0.33 min (16.33-16.67); 13%B for 3.33 min (16.67-20) | 0.3 | 60-900 | 20,000 @m/z 556 | 5 ppm | 5 ppm |  | ESI (P) | CID | 15-40 eV | UNIFI | MS, RT, MS/MS | UNIFI
18 | University of South Bohemia | Thermo Fisher Scientific | Qexactive HF | ThermoScientific, Hypersil BD | 150; 2.1; 2.4 |  | 5 | 50 | A=water+5mmol amonium acetate + 0.1% formic acid; B=acetonitrile+0.1% formic acid | as described in Appendix of manual for the study | as described in Appendix of manual for the study | 60-900 | 60000 FWHM @m/z 250; 40000 FWHM @m/z 800 | 2 ppm | 5 ppm |  | ESI (P) | HCD | as described in manual | CompoundDiscoverer 3.1 | MS, MS/MS, izotope patterns | Compound Discoverer embeded + own library, myCloud, Chemspider with using MassBank | MS, MS/MS, prediction of MS/MS, RT if available |  | MassFrontier
19 | Vitens | Sciex | Triple TOF 6600+ | Waters Xselect HSS C18 | 150; 4.6; 3.5 | ACN/Water 14/86 | 200 | 35 | A=water+5mM ammonium formetate; B=Acetonitril+0.1% formic acid | 13%B(0 min), 13%B (1 min), 50%B (2 min), 5%B (20.9 min), 5%B (23,8 min), 13%B (24,3 min), 13%B (29,2 min) | 1 | 77-1000 | 28,000 @m/z 130 43,000 @m/z 609 | 2 ppm | 5 ppm |  | ESI (P) |  | 35 eV | PeakView | MS, RT, MS/MS | Meta | MS, RT, MS/MS, prediction of retention time
20 | VITO | Thermo Fisher Scientific | Q Exactive | Acquity UPLC HSS C18 | 150; 2.1; 1.8 | direct injection | 5 | 50 | A=5mM NH4F in H2O (pH=3); B= 0.1%FA in Acetonitrile | 13 %B(0 - 0.5 min); 13 %B - 50%B over 9.5min; 50%B - 95%B(0.75); 95 %B (2 min); 13%B (0.25 min); hold 2.1 min 13% B | 0.4 | 60-900 | 70,000 @m/z 60; 70,000 @m/z 900 | 5 ppm | 5 ppm |  | ESI (P) | HCD | 30-40 eV | Trace Finder | MS, RT, MS/MS | Compound discoverer 3.1 | MS, RT, MS/MS, prediction of retention time, mzcloud | log Kow correlation | MetFrag | ug/L
21 | VU Amsterdam | Bruker Daltonics | Compact QToF model (Bruker, Germany) | Lunar ® Omega | 100; 2.1; 3 | was already solved | 5 | 50 | A=water+5 mM ammonium formate; B=ACN+0.1% formic acid | 13%B(0);13%B for 0.44 min;13%B-50% over 8.45 min;50%B-95% over 0.67 min;95%B(1.33);95%B-13% over 0.22 min; 13%B (2.22) | 0.3 | 60 - 900 | 25000 @m/z 1200 | 3 ppm | 5 ppm | 0.014 FWHM @ m/z 200 | ESI (P) | CID | 15 eV (@m/z 60), 25 eV (@m/z 500), 40 eV (@m/z 900) | - | - | MetaboScape® 4.0 | MS, MS/MS | - | - | -

## Feature Detection
ID_Institute | software package | feature detection method | mass resolution | peak width (s) | Intensity threshold | max number of peaks | compound score | Intensity threshold in high energy (MSE) | background filter | peak integration | xcmsSet ppm | xcmsSet prefilter | xcmsSet peakwidth | xcmsSet mzdiff | group method | group bandwith | group mzwid | groupminfrac | rt correction | Detect compound node: Max peak width (min) | Detect compound node: Mass tolerance (ppm) | Detect compound node: Intensity tolerance for isotope search (%) | Detect compound node: S/N | Detect compound node: min. peak intensity | Detect compound node: ions included | sample/blank threshold | Remove singlets | min # scans per peak | min # of isotopes | Mass tolerance | Isotope search intensity tolerance | Min. Peak intensity | S/N treshold | Sample/Blank ratio | Min. No. of isotopes | Max. peak width in half height | Minimum Peak Lenght (spectra) | Minimum Peak Lenght recursive (spectra) | Minimum Number of Features for extraction | Minimum Number of Features Present in Analysis | RT Range | Mass Range | EIC Correlation | Primary Ion | Seed Ions | Common Ions | Blank Substraction | CV(%) Technical replicates | Minimum # of mass spectral peaks | Total Intensity Threshold | S/N Threshold (FT-only) | min. peak intensity | min # scans per Peak | mass tolerance (ppm) | remove singlets | min. # Isotopes | Ions | Background | ppm | S/N | tR shift window for EIC selection | m/z diff | Gaussian fit | prefilter | Retention time correction | bwopt | minfracopt | gapInitopt | gapExtendopt | Calibration | area ratio threshold | minimum retention time | peak detection sensitivity | maximum retention | peak splitting | minimum peak width | s/n integration threshold | XIC width | gaussian smooth width | noise percentage | baseline subtract window | Lower RT limit (min) | Upper RT limit (min) | Min Prec. Mass (Da) | Max Prec. Mass (Da) | Min. Peak count | MS Order | Min Collision energy | Max Collision energy | Scan type | Polarity | Peak Filter (S/N) | Precurosr selection | Use Isotope Pattern in Precursor Reevaluation | Provide Profile Spetra | All other setttings | 1. Align retention time: Aligment mode | 1. Align retention time: Maximum shift (min) | 1. Align retention time: Mass tolerance | 2. Detect unknown compounds: Intensity tolerance for isotope search(%) | 2. Detect unknown compounds: S/N treshold for centroids | 2. Detect unknown compounds: Minimum peak intensity | 2. Detect unknown compounds: Ions | 2. Detect unknown compounds: Minimum element count for isotope pattern modeling | 2. Detect unknown compounds: Maximum element counts for isotope pattern modeling | 3. Group unknown compounds: mass tolerance | 3. Group unknown compounds: preferred ions | 4. Predict composition: mass tolerance | 4. Predict composition: Minimum element count for expected elements | 4. Predict composition: Maximum element counts for isotope pattern modeling | 4. Predict composition: minimum RBDE | 4. Predict composition: maximum RBDE | 4. Predict composition: minimun H/C | 4. Predict composition: maximum H/C | 4. Predict composition: maximum number of candidates | 4. Predict composition: Intensity tolerance | 4. Predict composition: Intensity treshold | 4. Predict composition: S/N ratio | 4. Predict composition: use dynamic recalibration | 4. Predict composition: use fragment matching | 4. Predict composition: mass tolerance for fragments | 4. Predict composition: S/N ratio for fragments
1 | Compound Discoverer | Select Spectra | 60000 | not specified | 1000000 | not specified | [75 empty] | 19 | 90 | 60 | 900 | 1 | MS2, MS1 | 0 | 1000 | Any | Is + | 1.5 | Use MS(n-1) Precursor | True | Automatic | Emtpy or Irrelevant
2 | XCMS + DeepLearner | CentWave + Peak Selection | 25000 | min (8.75)-max (42) | 300 | 30000 | [52 empty] | 18.75 | 3 | 60 sec | -0.001 | TRUE | 3-300 | Obiwarp + Loess | 5 | 0.5 | 0.928 | 2.292 | TRUE
2 | XCMS + DeepLearner | CentWave + Peak Selection | 25000 | min (8.25)-max (30) | 300 | 30000 | [52 empty] | 19 | 3 | 35 sec | -0.001 | TRUE | 3-300 | Obiwarp + Loess | 2.25 | 0.34 | 0.52 | 2.71 | TRUE
3 | UNIFI | 3D feature detection | automatic | automatic | 50 | 100000 |  |  | high | [31 empty] | 1 to 11 (only Norman) | 50-1200
4 | Digital sample freezing
5 | CompoundDiscoverer3.1 | CompoundDiscoverer3.1 | 70000 | workflow attached as txt
6 | CompoundDiscoverer3.1 | CompoundDiscoverer3.1 | MS1: 240K @ m/z 200; MS2: 15K @m/z 200 | 48 | 0 | [43 empty] | 2 | 0 | 1.5 | 50000 | 6 | 5 | TRUE | 2 | [2M+ACN+H]+1 [2M+ACN+Na]+1 [2M+H]+1 [2M+K]+1 [2M+Na]+1 [2M+NH4]+1 [M+2H]+2 [M+ACN+2H]+2 [M+ACN+H]+1 [M+ACN+Na]+1 [M+DMSO+H]+1 [M+H]+1 [M+H+K]+2 [M+H+MeOH]+1 [M+H+Na]+2 [M+H+NH4]+2 [M+H-H2O]+1 [M+H-NH3]+1 [M+K]+1 [M+Na]+1 [M+NH4]+1 | Sample/Blank < 5
6 | CompoundDiscoverer3.1 | CompoundDiscoverer3.1 | MS1: 240K @ m/z 200; MS2: 15K @m/z 200 | 18 | 0 | [43 empty] | 1 | 0 | 1.5 | 50000 | 3 | 5 | TRUE | 2 | [2M+ACN+H]+1 [2M+ACN+Na]+1 [2M+H]+1 [2M+K]+1 [2M+Na]+1 [2M+NH4]+1 [M+ACN+H]+1 [M+ACN+Na]+1 [M+H]+1 [M+H+MeOH]+1 [M+H-H2O]+1 [M+H-NH3]+1 [M+K]+1 [M+Na]+1 [M+NH4]+1 | Sample/Blank < 5
7 | CD discoverer 2.1 | Environmental with statistics unknown ID w online local databases searches | [94 empty] | Adaptive curve | 2 | 5 ppm | 30 | 3 | 100000 | [M+2H]2+, [M+Cl]-, [M+H]+, [M+Na]+, [M+NH4+]+, [M-e-]+, [M-H]-, [M-H+HAc]-1 | C, H | C90, H190, Br3, Cl4, F6, K2, N10, Na2, O18, P3, S5 | 5 ppm | [M+2H]2+, [M+Cl]-, [M+H]+, [M+Na]+, [M+NH4+]+, [M-e-]+, [M-H]-, [M-H+HAc]-1 | 5 ppm | C, H | C90, H190, Br3, Cl4, F6, K2, N10, Na2, O18, P3, S5 | 0 | 40 | 0.1 | 3.5 | 10 | 30 | 0.1 | 3 | TRUE | TRUE | 5 ppm | 3
8 | xcms |  |  | 20 | 700 | [53 empty] | 22.8
8 | xcms |  |  | 16 | 700 | [53 empty] | 9.2
9 | masshunter | molecular feature extractor | 10000 to 30000 |  | 15000 |  | 80
10 | UNIFI | 3D feature detection | 24000 | 6
10 | XCMS | centWave | [8 empty] | 15 | (5,50) | (3,15) | 0.007 | density | 8 | 0.01 | 0.65 | obiwarp
11 | Metaboscape | TREX |  |  | 1000 | [31 empty] | 6 |  |  |  | 0.4-18 | 60-900 | 0.8 | [M+H]+ | [M+K]+ [M+Na]+ [M+NH4]+ | [M-H2O+H]+
12 | UNIFI | 4D peak detection | automatic in UNIFI |  | 80 counts | 40000 |  | 50 counts
13 | R | patRoon package | 30000 | FWHM 1-24 | 5000 | 10000 |  |  |  |  | 5 | 39 | 1 | 9 | 0.004 | 2.5 | 0.005 | 0
14 | SciexOS | MQ4 algorithm | 30000 |  | 100 | [65 empty] | 3 |  |  |  | 2 points | 3 points | 3 | 0.02 Da | 0.8 points | 0.4 | 0.5 min
15 | Compound discoverer | Compound discoverer node based workflow | 60000 MS/AIF | [16 empty] | 0.5 | 5 | 30 | 3 | 100000 | [2M+H]+1; [2M+Na]+1; [M+2H]+2; [M+H]+1; [M+H+MeOH]+1; [M+H+Na]+2; [M+H+NH4]+2; [M+H-H2O]+1; [M+H-NH3]+1; [M+K]+1; [M+Na]+1; [M+NH4]+1; [M-e]+1 | 5 | yes | 8 | 1
16 | UNIFI | 3D feature detection |  | automatic | 250 | 300000 |  |  | high | 0.1-19.9 min
17 | UNIFI | 3D feature detection | 20000 | 6 | 100 | 10000 |  | 20
18 | Compound discoverer 3.1 | Environmental Unknown ID w Online and Local Database Searches_NIVA_02 |  |  | 500000 | [24 empty] | 5 ppm | 0.3 | 500000 | 3 | 10 | 2 | 0.5 min
19 | Peakview (Masterview) |  |  | 0.35 | 100 | no max
20 | Compound discoverer 3.1 | detect compounds | 70000 | not specified | 1000000 | not specified
21 | MetaboScape 4.0 | T-ReX | - | - | 2000 | - | [30 empty] | 10 | 8 | 2 | 2 | 0.31-27.96 min | 30-1300 m/z | 0.8 | [M+H]+ | [M+Na]+, [M+K]+ | [M-H2O+H]+ | Exclude features with intensities lower than 3x the intensity in the Blank | <20
21 | MetaboScape 4.0 | T-ReX | - | - | 2000 | - | [30 empty] | 8 | 6 | 2 | 2 | 0.31 - 13.3 min | 60 - 900 m/z | 0.8 | [M+H]+ | [M+Na]+, [M+K]+ | [M-H2O+H]+ | Exclude features with intensities lower than 3x the intensity in the Blank | <20